## HR
First author (year) | title | publication_year | publication_date | criterion | aw_model | metric | Population cohort | healthy | clinical_pop | total_participants | Males | Females | Procedure | free_living? | at_rest? | exercise?
Abt et al., (2017) | The validity and inter-device variability of the Apple Watch™ for measuring maximal heart rate | 2017 | 2017-11-01 | Polar T31 | First Generation | heart rate | Recreationally active participants | y | n | 15 | 8 | 7 | a single incremental maximal oxygen uptake test on a motorised treadmill | n | n | y
Alfonso et al., (2022) | Agreement between two photoplethysmography-based wearable devices for monitoring heart rate during different physical activity situations: a new analysis methodology | 2022 | 2022-09-14 | Biopac MP36 (ECG) | Series 6 | heart rate | University students | y | n | 18 | 10 | 8 | lying position for 5 min (lying activity), then to seat on a chair for 5 min (sitting activity), stand for 5 min (standing activity), and finally walk on a treadmill without inclination at a speed of 4 km/h for another 5 min (walking activity). | n | y | y
Al-Kaisey et al., (2019) | Accuracy of wrist-worn heart rate monitors for rate control assessment in atrial fibrillation | 2019 | 2019-11-18 | Holter monitor (DigiTrakXT, Philips, the Netherlands) | Series 3 | heart rate | Patients with AFib (26); patients with sinus rhy (6) | n | y | 20 | not reported | not reported | 24hour monitoring of HR. day: 7am-18:59; night: 19:00-6:59. | y | y | y
Bai et al., (2021) | Comprehensive comparison of Apple Watch and Fitbit monitors in a free-living setting | 2021 | 2021-05-26 | Polar H7 | Series 2 | heart rate | Healthy adults | y | n | 48 | 17 | 31 | 24hour free-living monitoring of HR, classified by 3 intensity levels: sedentary behaviour (SB), light physical activity (LPA), mod-vig (MVPA). (intensity level was determined by ActiGraph data estimation and METs | y | y | y
Behzadi et al., (2020) | Feasibility and Reliability of SmartWatch to Obtain 3-Lead Electrocardiogram Recordings | 2020 | 2020-09-07 | 12-lead ECG (Edan SE-1200 Express, CAmed Medical Systems GmbH, Köln, Germany) | Series 4 | heart rate | Patients in sinus rhythm referring to the Leipzig Heart Center, Germany | n | y | 100 | 62 | 38 | AW ECG taken immediately after standard ECG. HR obtained during this measurment too. rest. | n | y | n
Bent et al., (2020) | Investigating sources of inaccuracy in wearable optical heart rate sensors | 2020 | 2020-02-10 | ECG (Bittium Faros 180) | Series 4 | heart rate | not specified | not reported | not reported | 53 | 21 | 32 | Rest, deep breathing, activity (5 mins walking), typing | n | y | n
Düking et al., (2020) | Wrist-Worn Wearables for Monitoring Heart Rate and Energy Expenditure While Sitting or Performing Light-to-Vigorous Physical Activity: Validation Study | 2020 | 2020-05-06 | Polar H7 | Series 4 | heart rate | Healthy adults | y | n | 25 | 11 | 14 | two lab visits, 3 days between visits, 2 different wearables tested on each occasion, sitting, walking and running at different speeds (1.1 m/s, 1.9 m/s, 2.7 m/s, 3.6 m/s, and 4.1 m/s) for 5 minutes, interspersed with 5 minutes of standing still. All participants also performed 6 ~30-m sprints involving multiple changes in direction (ranging from 10° to 180°) | n | y | y
Etiwy et al., (2019) | Accuracy of wearable heart rate monitors in cardiac rehabilitation | 2019 | 2019-06-09 | ECG (Quinton Q-tel RMS telemetry system) | not reported | heart rate | Adults in phase 2 or 3 of cardiac rehab | n | y | 80 | 65 | 15 | Patients were assessed in two different exercise conditions: treadmill and stationary cycle. For both the treadmill and the stationary cycle, HR was measured at rest and at 3, 5, and 7 minutes of individually prescribed steady-state exercise at 50–70% of HR reserve. This provided HR measurements at 8 time points for each participant. | n | y | y
Falter et al., (2019) | Accuracy of apple watch measurements for heart rate and energy expenditure in patients with cardiovascular disease: Cross-sectional study | 2019 | 2019-03-19 | 12-lead ECG (Cardiosoft, General Electric Company) | First Generation | heart rate | Patients scheduled for a cardiopulmonary exercise test (CPET) as part of their cardiovascular rehabilitation program | n | y | 40 | 32 | 8 | Graded CPET on a cycle ergometer, HR analysed at 3 intensities: seated rest (HR1), moderate (HR2), maximal (HR3). The CPET started with 1 minute of seated rest. The exercise then started at 20 watts and load was increased with 20 W/min | n | y | y
Giggins et al., (2021) | Accuracy of Wrist-Worn Photoplethysmography Devices at Measuring Heart Rate in the Laboratory and During Free-Living Activities | 2021 | 2021-10-01 | Polar H10 | Series 5 | heart rate | Healthy adults | y | n | 10 | 5 | 5 | "Shapiro Wilk test- The treadmill protocol consisted of three five-minute stages; walking at 4km/h, jogging at 7km/h and running at 10km/h [10]. Each five minute stage of the protocol was interspersed with five minutes of rest sitting on a chair, and HR recovery was monitored for five minutes following the completion of the protocol" Real-world free-living validation was completed by one female (continuous HR recording for 12 hours) | n | y | y
Gillinov et al. (2017) | Variable Accuracy of Wearable Heart Rate Monitors during Aerobic Exercise | 2017 | 2017-08-01 | ECG (Quinton Q-tel RMS) | not reported | heart rate | Healthy adults | y | n | 50 | 23 | 27 | HR was measured when performing four different types of exercise at varying intensities; these included tread- mill, stationary bicycle, elliptical trainer with arm levers, and elliptical trainer without arm levers. The order of exercises was assigned randomly. //// all modes of exercise were performed at 3 different intensities (light/mod/vig) | n | n | y
Helmer et al., (2024) | Reliability of continuous vital sign monitoring in post-operative patients employing consumer-grade fitness trackers: A randomised pilot trial. | 2024 | 2024-04-11 | 3-lead ECG | Series 7 | heart rate | Post-op patients in hospital (nephrectomies, cystectomies and cystoprostatectomies, non-cardiac thoracic surgey). | n | y | 10 | 3 | 7 | presume a resting measurement: "The trackers then were used by the patients during their entire hospital stay. ... Moreover, all collected data points are time-stamped, which enables us to subsequently compare measurements taken at the same time by different devices." // Immediately post-op partiicpants were equipped with a fitness track that was randomly assigned. The wearable was used by the patients for their entire hospital stay. Researchers checked on the patients daily to ensure the wearable was positioned correctly. | n | y | n
Heyken et al., (2021) | COMPARISON OF WEARABLES FOR SELF-MONITORING OF HEART RATE IN CORONARY REHABILITATION PATIENTS | 2021 | 2021-06-01 | ECG (number of leads unreported) | Series 1 | heart rate | Patients in cardiac rehab sports groups who were in sinus rhythm | n | y | 35 | not reported | not reported | ECG-monitoring during cardiac rehabilitation training on a bicycle ergometer | n | n | y
Ho et al., (2022) | Accuracy of wrist-worn wearable devices for determining exercise intensity | 2022 | 2022-08-08 | 12-lead ECG (Custo cardio 100 BT) | Series 6 | heart rate | Healthy adults | y | n | 30 | 15 | 15 | subjects performed formal ramp incremental exercise tests on an electronically braked cycle ergometer equipped with a face mask. The exercise protocol started with the subjects resting on a seat for 3min, followed by a warmup with unloaded cycling at 0 W for 1 min, and a 15-W/min ramp protocol at a pedaling rate of 60–70 r/min. The tests terminated when the subjects were exhausted and unable to maintain 60 r/min. /// The exercise intensity was represented by the lower and upper limits of each exercise intensity, including 40% (lower limit of moderate intensity, represented as HR40), 60% (the cutoff point between moderate and vigorous intensity, represented as HR60), and 89% (upper limit of vigorous-intensity, represented as HR89). The target HR calculated from ECG was used as a reference for prescrib- ing exercise intensity. The target HRs from the two wrist- worn wearable devices were also calculated using the lower and upper limits of HRs obtained during moderate and vigorous-intensity exercises. In addition, age-predicted HRmax formulas, including Fox and Tanaka equations, were used to establish exercise intensity using the HRR for- mulas for medium- and vigorous-intensity exercises. | n | n | y
Huynh et al., (2021) | Heart rate measurements in patients with obstructive sleep apnea and atrial fibrillation: Prospective pilot study assessing Apple Watch’s agreement with telemetry data | 2021 | 2021-02-08 | Telemetry (CARESCAPE Monitor B650; GE Healthcare) [likely 12-lead ECG] | First Generation | heart rate | Patients with obstructive sleep apnea in AFib | n | y | 20 | 17 | 3 | Sleep: The Apple Watch face and telemetry monitor (CARESCAPE Monitor B650; GE Healthcare) were observed concomitantly under video recording in the presence of a study co-investigator (RS) for 90 seconds. After excluding the first 30 seconds of data to allow time for the watch’s HR monitor to equilibrate, HR measurements were sampled every 10 seconds for 60 seconds, yielding a total of 7 observations per participant per device (Apple Watch and telemetry). | n | y | n
Hwang et al., (2019) | Assessing accuracy of wrist-worn wearable devices in measurement of paroxysmal supraventricular tachycardia heart rate | 2019 | 2019-05-01 | 12-lead ECG (Prucka, GE Medical Systems Information Technology, Milwaukee, WI, USA) | Series 2 | heart rate | Patients with history of paroxysmal palpitations or documented paroxysmal supraventricular tachyarrhythmia (SVT) | n | y | 51 | 27 | 24 | An accurate measurement was defined when a difference between the device-measured HR and an ECG-recorded HR is 5 beats per minute (bpm) or less. The HR was measured in the following situations: 1) Baseline HR prior to the EPS 2) Induced SVT HR 3) Atrial pacing study after successful ablation: A. pacing cycle lengths (PCLs): 600 ms, 500 ms, 430 ms, 380 ms, 330 ms, 300 ms B. performed only 1: 1 atrioventricular (AV) conduction was maintained C. at least a 30-second resting period between pacing studies 4) Ventricular pacing study after successful ablation: A. PCLs: 600 ms, 500 ms, 430 ms, 380 ms, 330 ms, 300 ms B. at least a 30-second resting period between pacing studies | n | y | n
Khushhal et al., (2025) | Accuracy of Apple Watch to Measure Cardiovascular Indices in Patients with Cardiac Diseases: Observational Study | 2025 | 2025-09-02 | Polar H10 | Series 8 | heart rate | 260 male patients with cardiac diseases (mean [SD]; age 47 (15) years), including 190 ischemic patients with regular heart rhythm and 70 patients with irregular rhythm | n | y | 260 | 260 | 0 | measure the heart rate at rest for 5 minutes, followed by moderate intensity exercise [50–70% Heart rate reserve (HRR)], 50 to 70% of HRR was defined as moderate exercise intensity,23 for 16 minutes, and 3 recovery minutes after the exercise session. // "The exercise session was about 16 minutes in duration at moderate intensity according to the HRR using the cycle ergometer, preceded by a 3-minute warm-up, and followed by a 3-minute cool-down. The HR was recorded every 30 seconds during the whole period, and the oxygen saturation was recorded five times pre-and-post exercise." | n | n | y
Khushhal et al., (2025) | Accuracy of Apple Watch in Measuring 30-Second Resting Electrocardiography in Patients with Cardiac Diseases and Comorbidity: An Observational Cross-Sectional Study. | 2025 | 2025-01-30 | 12-lead ECG (ECG300G, Contec Medical Systems Co., Ltd., China) | Series 8 | heart rate | 34 cardiac patients with sinus rhythm, and 14 cardiac patients with irregular HR rhythm. Patients with chronic diseases, including 14 diabetics, 34 hypertension patients, and 16 dyslipidemia patients | n | y | 112 | 112 | 0 | simultaneous ECG and Apple Watch measurement at rest | n | y | n
Khushhal et al., (2017) | Validity and Reliability of the Apple Watch for Measuring Heart Rate During Exercise. | 2017 | 2017-09-16 | Polar S810i monitor | First Generation | heart rate | Healthy males able to complete 3x bouts of 11 minute exercise | y | n | 21 | 21 | 0 | Exercise involved 5-min bouts of walking, jogging, and running at speeds of 4km.h−1, 7km.h−1, and 10km.h−1, followed by 11 min of rest between bouts (on treadmill). One apple watch was placed on each wrist - one left, one right. inter and intra-device reliability reported. | n | y | y
Kim et al., (2023) | Validation of Wearable Digital Devices for Heart Rate Measurement During Exercise Test in Patients With Coronary Artery Disease | 2023 | 2023-06-22 | 12 lead ECG | Series 7 | heart rate | coronary artery disease patients | n | y | 44 | 36 | 8 | modified Bruce protocol // Cardiopulmonary exercise test via modified Bruce protocol with (Quark CPET; COSMED). Heart rate (HR) measurement during cardiopulmonary exercise test. HR was recorded at 2, 4, and 6 minutes in rest phase. HR was recorded every minute beginning from 1 minute during exercise phase until termination of the test. HR at termination, which was the beginning of recovery phase, was also recorded and once every minute subsequently in recovery phase. | n | y | y
Koshy et al., (2018) | Smart watches for heart rate assessment in atrial arrhythmias | 2018 | 2018-02-19 | ECG (IntelliVue, Phillips, The Netherlands) | Series 1 | heart rate | Patients from coronary care unit with and without arrhythmia a) Atrial arrhythmias and b) Sinus Rhythm // patients from coronary care unit, intensive care unit and emergency room | n | y | 102 | 66 | 48 | Supine measurement | n | y | n
Lee et al., (2022) | COMPARISON OF APPLE WATCH SERIES 4 VS. KARDIAMOBILE: A TALE OF TWO DEVICES | 2022 | 2022-07-18 | 12-lead ECG | Series 4 | heart rate | Patients from the Toronto Heart Centre cardiology clinic | n | y | 200 | 118 | 82 | HR readings were considered correct if they deviated a maximum of  5 beats per minute from the 12-lead ECG, taking into account variability in atrial fibrillation (AF). those who signed the informed consent form underwent their usual scheduled 12-lead ECG. Immediately afterward, they were asked to sit down and were properly fitted with both the AW4 and the KM for the study duration. The AW4 was strapped onto the wrist of their nondominant hand, in the typical position of a wristwatch, ensuring that it rested snugly against their skin. The participants were instructed to use the ECG and HR functions on the AW4, taking 2 sets of 30- second readings for each mode. | n | y | n
Montalvo et al., (2023) | Commercial Smart Watches and Heart Rate Monitors: A Concurrent Validity Analysis | 2023 | 2023-09-01 | 12-lead ECG (Quinton Q-Stress Cardiac Stress System, Hill-Rom., IL) note Polar H10 for exercise | Series 6 & 7 | heart rate | Physically active young adults 18-30yrs // Twenty-four male collegiate football players and 20 college-age recreationally active young adults | y | n | 43 | 33 | 10 | a treadmill- based exercise session. The testing protocol included 3 minutes of standing still (resting), then walking at low intensity, jogging at a moderate intensity, running at a high intensity, and postexercise recovery. subjects were required to be free of any caffeine intake for 12 hours, to continue with their regular nutrition, and be well hydrated before the testing session. All sessions were conducted at the same time for all subjects (8:00 AM–12:00 PM). resting heart rate was recorded for 3 minutes while standing still. subjects were asked to walk (low-intensity), jog (moderate-intensity), and run (high-intensity) on the treadmill at zero % incline. Each stage was 3 minutes in duration. Football athletes performed these exercise intervals at 3, 6, and 9 mph, respectively. Recreationally trained subjects followed the same protocol except the speed set on the treadmill was lowered to walk at 2.5 mph, jog at 5 mph, and run at 7.5 mph. | n | y | y
Mulholland et al., (2025) | Influence of skin pigmentation on the accuracy and data quality of photoplethysmographic heart rate measurement during exercise. | 2025 | 2025-09-01 | Polar H10 | Series 8 | heart rate | 28 healthy adults (White, n = 16; Black, n = 10; Middle Eastern, n = 1; Persian, n = 1) | y | n | 28 | 15 | 13 | recorded continuously during rest, cycling, and recovery; HR data were averaged into 30-s epochs for analysis (results give pooled values, not broken down as rest vs exercise) | n | n | y
Nelson et al., (2019) | Accuracy of consumer wearable heart rate measurement during an ecologically valid 24-hour period: Intraindividual validation study | 2019 | 2019-03-11 | 3-lead ambulatory ECG (Vrije Universiteit Ambulatory Monitoring System) | Series 3 | heart rate | Caucasian male | y | n | 1 | 1 | 0 | 24 hour recording which included 5 different conditions: sitting, walking, treadmill running, ADLs, sleeping. | y | y | y
Nuss et al., (2019) | Accuracy of Heart Rate and Energy Expenditure Estimations of Wrist-Worn and Arm-Worn Apple Watches | 2019 | 2019 | 12-lead electrocardiograph (ECG; Q-Stress ECG, Mortara, Milwaukee, WI, USA) | Series 1 | heart rate | Young adults | y | n | 30 | 15 | 15 | The exercise protocol started with the participant sitting in a chair for three minutes, followed by standing on the treadmill for two minutes as HR readings were taken from each device every minute. The lowest ECG reported HR from the rest period was recorded as the participant’s seated HR. After this ve-minute rest period, the participant started the Bruce protocol. // Research staff recorded HR estimations from each device and from the ECG at the end of each minute of the Bruce protocol. | n | y | y
O'Grady et al., (2024) | The Validity of Apple Watch Series 9 and Ultra 2 for Serial Measurements of Heart Rate Variability and Resting Heart Rate. | 2024 | 2024-09-26 | Polar H10 + Kubios | Series 9 and Ultra 2 | heart rate | healthy | y | n | 39 | 17 | 22 | RHR // each morning upon waking, participants were instructed to fit their Apple Watch and Polar H10 according to the manufacturer’s instructions. They simultaneously recorded 5-min HRV readings using both devices. HRV data from the Polar H10 were captured using the Kubios HRV mobile app, while Apple Watch data were recorded via the Breathe app on Apple Watch. Measurements were taken at the same time each morning under consistent, relaxed environmental conditions, with participants lying supine and keeping still. | y | y | n
Pasadyn et al., (2019) | Accuracy of commercially available heart rate monitors in athletes: A prospective study | 2019 | 2019-08-09 | 3-lead ECG | Series 3 | heart rate | Healthy athletic adults (able to run a mile in under 7 mins) | y | n | 50 | 34 | 16 | running at 4, 5, 6, 7, 8, and 9 mph at zero incline on a treadmill // Subjects were asked to run for at least 2 min at each of these speeds. HR was assessed from the four devices at 2 min of activity | n | n | y
Reece et al., (2021) | Assessing Heart Rate Using Consumer Technology Association Standards | 2021 | 2021-06-30 | Polar H7 [The PerformTek app (Valencell, Inc, Raleigh, NC, USA) collected data from the Polar chest strap] | Series 4 | heart rate | Physically active adults that passes the PARQ+ | y | n | 23 | 13 | 10 | sitting, activities of daily living, walking, jogging, running, and stationary cycling, totaling 57 min of monitored activity. //// sedentary, light, mod, vigorous intensity activities defined by METs/RPE or % of HRmax | y | n | y
Saghir et al., (2020) | A comparison of manual electrocardiographic interval and waveform analysis in lead 1 of 12-lead ECG and Apple Watch ECG: A validation study | 2020 | 2020-08-26 | 12-lead ECG [Philips Page- Writer TC70 ECG machine (Philips USA, Andover, MA)] + analysis by 3 internal medicine residents | Series 4 | heart rate | Healthy adults | y | n | 43 | 15 | 28 | The AW ECG and standard 12-lead ECG were performed consecutively, with patients given instruc- tions to lie as still as possible without shifting positions for the entire duration of both recordings. Due to signal interference, the AW ECG and 12-lead ECG could not be performed simultaneously. The 12-lead ECG was always obtained first, followed by disconnection of the lead wires and immediate placement of the AW on the patient’s left wrist to record the single-lead AW ECG. Time between both ECGs per participant typically was <60 seconds. | n | y | n
Sequeira et al., (2020) | Common wearable devices demonstrate variable accuracy in measuring heart rate during supraventricular tachycardia | 2020 | 2020 | 12-lead electrocardiogram | First Generation | heart rate | people with documented reentrant SVT | n | y | 24 | 11 | 13 | rest with electical stimulation : The diagnostic electrophysiological procedure for all participants proceeded per the regular standard of care. This included atrial and ventricular stimulation to induce multiple episodes of the clinical SVT before the planned ablation. Stimulation included premature extrastimuli and trains of burst pacing. A 12-lead electrocardiogram (ECG) was continuously recorded. | n | n | n
Seshadri et al., (2019) | Accuracy of the apple watch 4 to measure heart rate in patients with atrial fibrillation | 2019 | 2019-12-13 | 6-lead telemetry // continuous six-lead telemetry according to standard clinical practice on the step-down floor | Series 4 | heart rate | Post-op cardiac surgery patients, 50% in SR, 50% in AF | n | y | 50 | 36 | 14 | Each enrolled patient had a minimum of three assessments of heart rate per day for at least two days, generating a minimum of six data points per patient. [presumably at rest, but not explicitly stated] | n | y | n
Støve et al., (2023) | Accuracy of the Apple Watch Series 6 and the Whoop Band 3.0 for assessing heart rate during resistance exercises | 2023 | 2023-02-19 | Polar H10 | Series 6 | heart rate | Healthy adults | y | n | 29 | 13 | 16 | Participants completed five resistance exercises; Barbell Back Squat, Barbell Deadlift, Dumbbell Curl to Overhead Press, Seated Cable Row, and Burpees. During the exercises, heart rate was concurrently measured using the Polar H10, the Apple Watch Series 6. Before measurements, each participant was instructed to cycle on a cycle ergometer (Corival V2, Lode BV, Groningen, The Netherlands) at a load of 100 W for a five- minute warm-up period. Upon completing the warm-up, the participants transitioned to a resistance exercise protocol. Each exercise consisted of three sets of 15 repetitions. | n | n | y
Støve et al., (2020) | Measurement latency significantly contributes to reduced heart rate measurement accuracy in wearable devices | 2020 | 2020-05-14 | Polar H10 | Series 2 | heart rate | Healthy adults | y | n | 30 | 15 | 15 | Participants completed three submaximal exercise conditions; cycling (at 3 intensities), treadmill walking/running (various intensities) and rapid arm movements. During the exercises, HR was concurrently measured using the Polar H10 (PL), the Apple Watch 2 (AW) //// **Based on previous findings a criterion of precision of ±3 beats per minute [26] was then elected to deter- mine when the HR measures from the AW and GF were equivalent to that of the criterion measures. | n | n | y
Thomas et al., (2022) | Validity of Optical Heart Rate Measurement in Commercially Available Wearable Fitness Tracking Devices | 2022 | 2022-09-30 | Polar H10 | Series 4 | heart rate | healthy Adults | y | n | 20 | 10 | 10 | 4 different conditions: sedentary (SED), treadmill walking (WLK), running (RUN), and dynamic running/walking (DYN). SED, WLK, and RUN were all 7-minute protocols with 1 minute of standing, 5 minutes of prescribed activity, and 1 final minute of standing. The DYN protocol was a 12- minute protocol with 1 minute of standing, 10 minutes of variable intensity walking and running, and 1 minute standing. | n | y | y
Thomson et al., (2019) | Heart rate measures from the Apple Watch, Fitbit Charge HR 2, and electrocardiogram across different exercise intensities | 2019 | 2019-01-18 | 12-lead ECG (Q-Stress ECG, Mortara, Milwaukee, WI, USA) | First Generation | heart rate | Young healthy adults | y | n | 30 | 15 | 15 | Bruce treadmill protocol. HR readings were taken from each device and from the ECG each minute for the entire duration of the exercise protocol. The exercise protocol began with three min- utes of upright sitting in a chair, followed by two minutes of standing on the treadmill. //// Intensity categories were very light (<20% HRR), light (20–40% HRR), moderate (40–60% HRR), vigorous (60–85% HRR), and very vigorous (>85% HRR) (American College of Sports Medicine, 2013). /// equivalence set at 5% | n | y | y
Uphill et al., (2024) | Validity of Apple Watch, Garmin Forerunner® 935 and GENEActiv for estimating energy expenditure during close quarter battle training in Special Forces soldiers | 2024 | 2024-03-18 | Polar Team2 Pro | Series 5 | heart rate | Male soldiers | y | n | 49 | 49 | 0 | Close quarters combat training activity // The scenario involved carrying rifles, fast roping from a mock helicopter, walking and running and engaging targets as required. This activity takes at least 20 minutes to complete. | n | n | y
Wallen et al., (2016) | Accuracy of heart rate watches: Implications for weight management | 2016 | 2016-05-27 | 3-lead ECG (CASE exercise testing system, GE Healthcare, UK) | First Generation | heart rate | Healthy adults | y | n | 22 | 11 | 11 | supine and seated rest, walking and running on a treadmill and cycling on an ergometer // 1-hr proto- cols involving supine and seated rest, walking and running on a treadmill and cycling on an ergometer. /// Participants initially performed five-minute periods of supine, sit- ting and standing, respectively. Three stages of a Bruce graded treadmill exercise protocol were then undertaken followed by five minutes of seated rest. Participants then completed six, three- minute stages of a 25-watt step test (commencing at 25 W) on a cycle ergometer followed by a final five minutes of seated rest. | n | y | y

## AFib
First author (year) | title | publication_year | publication_date | criterion | aw_model | metric | Population cohort | healthy | clinical_pop | total_participants | Males | Females | Procedure | free_living?
Abu-Alrub et al., (2022) | Smartwatch Electrocardiograms for Automated and Manual Diagnosis of Atrial Fibrillation: A Comparative Analysis of Three Models | 2022 | 2022-02-04 | 12L ECG + Cardiac electrophysiologists | Series 5 | afib | 100 patients in AF and 100 patients in sinus rhythm |  |  | 200 | 112 | 88 | 30s ECG tracings with AW directly after ECG // Rest | n
Apple (2020) | Using Apple Watch for Arrhythmia Detection | 2020 | decembed_2020 | 12-lead ECG, 2 blinded cardiologists, disagreements solved by a third cardiologist | not reported | afib | Study subjects included those with normal SR at the time of screening—with no known history of AFib—and those with known persistent, permanent, or chronic AFib who were in AFib during screening | y | y | 546 | not reported | not reported |  | n
Briosa E Gala et al., (2025) | Diagnostic performance of single-lead electrocardiograms from a smartwatch and a smartring for cardiac arrhythmia detection | 2025 | 2025-03-26 | 12-lead ECG, 2 cardiologists, disagreements solved by a third cardiologist | Series 6 | afib | Individuals with cardiovascular disease |  |  | 483 | 344 | 139 | Simultaneous 12-lead ECG and AW 30s recordings | n
Ford et al., (2022) | SMART WARS - COMPARISON OF THE APPLE WATCH SERIES 4 AND KARDIABAND SMART WATCH TECHNOLOGY FOR THE DIAGNOSIS OF ATRIAL FIBRILLATION | 2022 | 2022-02-27 | 12-lead ECG | Series 4 | afib | Patients from outpatient cardiology clinics 65yrs or older |  |  | 125 | 62 | 63 | ECG recording, then AW recording on left wrist // patient lying supine | n
Lee et al., (2022) | COMPARISON OF APPLE WATCH SERIES 4 VS. KARDIAMOBILE: A TALE OF TWO DEVICES | 2022 | 2022-07-18 | 12-lead ECG | Series 4 | afib | recruited from the Toronto Heart Centre, a community-based cardiology clinic |  |  | 200 | 118 | 82 | They had a 12-lead electrocardiogram (ECG), followed by wearing the AW4 watch (OS 5.3), and pressing on the KM electrode plates, within the span of 5 minutes of one another. Each session involved a 12-lead ECG, an ECG from each device, and AW4’s photoplethysmography function (APPG). | n
Mannhart et al., (2022) | Clinical validation of an artificial intelligence algorithm offering cross-platform detection of atrial fibrillation using smart device electrocardiograms | 2023 | 2023-04-29 | 12-lead ECG (Schiller SDS-200, Touch 4.4.3; Schiller, Baar, Switzerland) + cardiologist interpretation | Series 6 | afib | Patients presenting at University Hospital of Basel (83% with know Hx of AFib) // Patients at routine cardiac follow up appointment |  |  | 117 | 85 | 32 | Patients eligible for participation obtained a preprocedural 12-lead electrocardiogram (part of routine clinical care), followed sharply by the five 30-second wearable smart device recordings. | n
Mannhart et al., (2023) | Clinical Validation of 5 Direct-to-Consumer Wearable Smart Devices to Detect Atrial Fibrillation: BASEL Wearable Study | 2023 | 2023-02-01 | 12-lead ECG (Schiller SDS – 200, Touch 4.4.3, Schiller) + cardiologist interpretation | Series 6 | afib | Patients presenting at University Hospital Basel for cardiac procedures. 30% patients had AFib. // Patients scheduled for catheter ablation procedures, electric cardioversions, pacemaker, or implantable cardioverter-defibrillator implantation were included. |  |  | 201 | 139 | 62 | Once enrolled, patients un- derwent a preprocedural 12-lead ECG (part of routine clinical care) followed by 5 wearable smart device recordings for 30 seconds (Apple Watch 6, AliveCor KardiaMobile or KardiaMobile 6L, Fitbit Sense, Samsung Galaxy Watch 3, and Withings Scanwatch). | n
Muller et al., (2024) | Validity of a smartwatch for detecting atrial fibrillation in patients after heart valve surgery: a prospective observational study | 2024 | 2024-05-25 | 6-lead telemetry - hospital telemetry system (a six-lead ECG, BeneVision TM80, Bio-Medical Electronics Co., Mindray, Shenzhen, China) // All participants were continuously monitored with the hospital telemetry system from the first to thee fifth postoperative day in hospital settings // A 12-lead ECG was taken with a Mortara XScribe 6, a standard 12-lead ECG machine (Scan-Med a/s, Drammen, Norway) during medical examinations 3days after heart valve surgery. | Series 5 | afib | Post-heart valve surgery // Post op left sided open heart valve surgery |  |  | 93 | 73 | 20 | All patients underwent single-lead ECG using the smartwatch three times or more daily on the second to third or third to fourth postoperative day. These results were compared with continuous 2–4 days ECG telemetry monitoring and a 12-lead ECG on the third postoperative day. // the smartwatch results were compared with the physician's evaluation of continuous telemetry and a 12-lead ECG. // By using an Apple Watch Series 5, single-lead ECGs were recorded up to three times a day on the second and third, or third and fourth postoperative days in the hospital. The ECG measurements were taken at specific times (morning, afternoon, and evening) during the day. In accordance with the smartwatch settings to ensure quality and accuracy, the ECG measurements were of 30 s duration. Additionally, patients were encouraged to take an extra ECG in case of clinical symptoms of AF such as palpitations, chest pain, dizziness, or heavy breathing. // 3 days post op a 12-lead ECG was taken with a Mortara XScribe 6, a standard 12-lead ECG machine (Scan-Med a/s, Drammen, Norway). The smartwatch single-lead ECG was taken within 10 min after the 12-lead ECG | n
Pepplinkhuizen et al., (2022) | Accuracy and clinical relevance of the single-lead Apple Watch electrocardiogram to identify atrial fibrillation | 2022 | 2022-12-15 | 12-lead ECG + 2 physicians | Series 6 | afib | Patients scheduled for elective electrical cardioversion (ECV) for AF in our center between February and June 2021 |  |  | 74 | 59 | 15 | AW was positioned on the left wrist and patients were in- structed to rest their arm on the table or their legs and keep their arms still during the recording. A test run for correct posi- tioning of the finger was done by each patient before the actual measurements and a member of the research team was present to assure proper measurements. The AW ECG was obtained during AF before ECV and during SR after ECV. If ECV was unsuccessful and SR was not achieved, no AW ECG post-ECV was obtained. // In case of an inconclusive or poor recording notification, the AW ECG recording was obtained up to 3 times. | n
Perez et al., (2019) | Large-scale assessment of a smartwatch to identify atrial fibrillation | 2019 | 2019-11-14 | ePatch (BioTelemetry, Philips) ECG patch + 2 clinicians + a third for disagreements | Multiple AW series included (Series 1 to Series 4) | afib | Adults 22yrs+ without AFib with an iPhone and Apple Watch | y | y | 419297 | 238700 | 177087 | Participants without atrial fibrillation (as reported by the participants themselves) used a smartphone (Apple iPhone) app to consent to monitoring. If a smartwatch- based irregular pulse notification algorithm identified possible atrial fibrillation, a telemedicine visit was initiated and an electrocardiography (ECG) patch was mailed to the participant, to be worn for up to 7 days. Surveys were administered 90 days after notification of the irregular pulse and at the end of the study. The main objectives were to estimate the proportion of notified participants with atrial fibrillation shown on an ECG patch and the positive predictive value of irregular pulse intervals with a targeted confidence interval width of 0.10. | y
Racine et al., (2022) | Role of Coexisting ECG Anomalies in the Accuracy of Smartwatch ECG Detection of Atrial Fibrillation | 2022 | 2022-10-12 | 12-lead ECG | Series 5 | afib | Hospital patients |  |  | 734 | 426 | 308 | All patients underwent 12-lead ECG directly followed by a 30-second ECG tracing with the use of an Apple Watch series 5, after receiving standardised instructions. The smartwatch’s automated AF detection algorithm then concluded “no signs of AF,” “AF,” or “not checked for AF (unclassified).” Unclassified recordings were “low heart rate” (below 50 beats/min), “high heart rate” (above 150 beats/min), “poor recording,” or “inconclusive recording.” | n
Saghir et al., (2022) | CORRELATION OF ATRIAL FIBRILLATION DETECTION USING OURA RING WITH PHOTOPLETHYSMOGRAPHY IN COMPARISON TO THE APPLE WATCH ELECTROCARDIOGRAPHY ALGORITHM (DH-576-04) | 2022 | 2022-05-01 | 3-lead ECG | Series 4 | afib | Patients who have undergone cardioversion or AF ablation to SR |  |  | 18 | 16 | 2 | AW ECG alongside standard ECG data were collected for 30 seconds. All device’s datasets were collected both before and after cardioversion or AF ablation. ECG diagnoses were confirmed by an Electrophysiologist. | n
Scholten et al., (2021) | A comparison of over-the-counter available smartwatches and devices for electrocardiogram based detection of atrial fibrillation | 2021 | 2021-12-29 | 12-lead ECG + cardiologist | Series 5 | afib | Patients with AF presenting for electrical cardioversion (ECV) |  |  | 220 | 143 | 77 | *** uninterpretable ECGs were excluded** We collected a standard 12-lead ECG recording immediately followed by four times 30 seconds of ECG recordings from different devices for every patient prior to the ECV. These paired measurements were considered simultaneous. | n
Seshadri et al., (2020) | Accuracy of Apple Watch for Detection of Atrial Fibrillation | 2020 | 25/02/2020 | 6-lead telemetry | Series 4 | afib | Post-op cardiac surgery patients |  |  | 50 | 36 | 14 | Rhythm assessments were performed 3 times per day (morning, early/late afternoon, and early/late evening) over 2 days, resulting in 6 assessments per patient. At each assess- ment, the subject was randomly assigned by computer program 1 of 5 AW4 watches that was placed on a wrist (randomly assigned); 5 watches were purchased to mitigate the impact of possible hardware malfunctions and the potential impact of a single device. When instructed, subjects depressed the crown for 30 seconds, and the watch reading (SR, AF, inconclusive) was recorded. Simultaneously, a PDF of the waveform on the Apple Health App and a rhythm strip from telemetry were saved and adjudicated by a cardiologist blinded to the patient’s telemetry rhythm. | n
Velraeds et al., (2023) | Improving Automatic Smartwatch Electrocardiogram Diagnosis of Atrial Fibrillation by Identifying Regularity within Irregularity | 2023 | 2023-11-20 | 12-lead ECG | Series 5 | afib | Patients with and without cardiac disease (24% without, 21% diagnosed with AF) // patients hospitalized in the Cardiology department |  |  | 144 | not reported | not reported | At rest conditions, participants underwent a standard 12-lead electrocardiogram, followed by a smartwatch ECG recording using the Apple Watch Series 5. /// Inconclusive results were considered as false results. The first completed smartwatch ECG tracing was used | n
Wasserlauf et al., (2023) | Accuracy of the Apple watch for detection of AF: A multicenter experience | 2023 | 21/03/2023 | Insertable cardiac monitor (ICM) or cardiac implanted electronic device (CIED) | Series 5 | afib | Subjects with a history of nonpermanent AF and either an insertable cardiac monitor (ICM) or cardiac implanted electronic device (CIED) with <5% ventricular pacing |  |  | 30 | 18 | 12 | Subjects with a history of nonpermanent AF and either an insertable cardiac monitor (ICM) or cardiac implanted electronic device (CIED) with <5% ventricular pacing were fitted with an AW Series 5 for 6 months and asked to wear for at least 14 hours per day. AF episodes were compared between the ICM/CIED and IRN. The primary endpoints were sensitivity, specificity, positive predictive value (PPV), and negative predictive value (NPV) of the IRN by subject for AF ≥1 h. Secondary endpoints were sensitivity and PPV by AF episode ≥1 h. Analysis was limited to a maximum of 10 ICM/CIED episodes per subject and included only those AF episodes occurring during active AW use confirmed by activity data. | y
Wouters et al., (2025) | Comparative Evaluation of Consumer Wearable Devices for Atrial Fibrillation Detection: Validation Study. | 2025 | 2025-01-09 | 12-lead ECG, MAC5000, GE Healthcare, Chicago, IL | not reported | afib | Patients visting the outpatient cardiology clinic and patients who were scheduled for electrical cardioversion at the clinic. |  |  | 122 | 78 | 44 | After the participants provided informed consent, they were instructed by study personnel to perform heart rhythm measurements using 4 CWDs in a random sequence consisting of 2 ECG- and 2 PPG-based devices, according to the manufacturer’s instructions. During all measurements, the patients were also instructed to stay still to avoid motion artifacts. // If good-quality measurements could not be obtained with a specific device, the patient’s results were not excluded from the analyses. | n

## SpO2
First author (year) | title | publication_year | publication_date | criterion | aw_model | metric | Population cohort | healthy | clinical_pop | total_participants | Males | Females | Procedure | free_living?
Arslan et al., (2023) | Accuracy of the Apple Watch in measuring oxygen saturation: comparison with pulse oximetry and ABG | 2023 | 2023-07-13 | Arterial blood gases used as reference standard. AW also compared to Masimo Rainbow Set pulse oximeter (Masimo, Irvine, CA, USA). | Series 6, 40mm | SpO2 | Patients with acute exacerbation of COPD | n | y | 167 | 99 | 68 |  | n
Apple (2022) | Performance Accuracy: Blood Oxygen app on Apple Watch | 2022 | October 2022 | ABG | Series 6 | SpO2 | Healthy adults | y | n | 24 | not reported | not reported |  | n
Helmer et al., (2024) | Reliability of continuous vital sign monitoring in post-operative patients employing consumer-grade fitness trackers: A randomised pilot trial. | 2024 | 2024-04-11 | FAST Sensor M1191B, Philips Healthcare, [checked – it is CE certified] 'Transmissive pulse oximetry' | Series 7 | SpO2 | Post-op patients in hospital | n | y | 10 | 3 | 7 |  | n
Jiang et al., (2023) | Investigating the accuracy of blood oxygen saturation measurements in common consumer smartwatches. | 2023 | 2023-07-12 | (Masimo MightySat Rx SpO2) Clinical-grade pulse oximeter | Series 7 | SpO2 | In or outpatients, majority with disease causing hypoxemia // Two participants did not have any disease-causing hypoxemia, while 42 had pulmonary disease, three had cardiovascular disease, and two had both pulmonary and cardiovascular disease | n | y | 49 | 31 | 18 |  | n
Jiang et al., (2024) | Performance of Wearable Pulse Oximetry During Controlled Hypoxia Induction | 2024 | 2024-07-16 | arterial blood oxygen saturation readings obtained from arterial blood gas samples. | Series 7 | SpO2 | Healthy participants | y | n | 9 | 5 | 4 |  | n
Khushhal et al., (2025) | Accuracy of Apple Watch to Measure Cardiovascular Indices in Patients with Cardiac Diseases: Observational Study | 2025 | 2025-09-02 | CMS50DL, Contec Medical Systems Co., Ltd., China | Series 8 | SpO2 | 260 male patients with cardiac diseases (mean [SD]; age 47 (15) years), including 190 ischemic patients with regular heart rhythm and 70 patients with irregular rhythm | n | y | 260 | 260 | 0 | measure the heart rate at rest for 5 minutes, followed by moderate intensity exercise [50–70% Heart rate reserve (HRR)], 50 to 70% of HRR was defined as moderate exercise intensity,23 for 16 minutes, and 3 recovery minutes after the exercise session. // "The exercise session was about 16 minutes in duration at moderate intensity according to the HRR using the cycle ergometer, preceded by a 3-minute warm-up, and followed by a 3-minute cool-down. The HR was recorded every 30 seconds during the whole period, and the oxygen saturation was recorded five times pre-and-post exercise." | n
Rafl et al., (2022) | Commercial smartwatch with pulse oximeter detects short-time hypoxemia as well as standard medical-grade device: Validation study | 2022 | 2022-10-11 | Masimo Radical-7 (pulse oximeter) | Series 6 | SpO2 | Healthy adults | y | n | 24 | 19 | 5 |  | n
Rajakariar et al., (2024) | Accuracy of Smartwatch Pulse Oximetry Measurements in Hospitalized Patients With Coronavirus Disease 2019 | 2024 | 2024-03-01 | Welch Allyn Connex Spot Monitor | Series 7 | SpO2 | patients hospitalized with COVID-19 infection, not requiring intensive care // 24% of patients were hypoxic at the time of assessment (n=48), and in this cohort, 77% were receiving supplemental oxygen (nasal oxygen, n=27; high flow oxygen, n=9; noninvasive ventilation, n=1). All other patients (n=163) were breathing room air when their readings were obtained. | n | y | 200 | 108 | 92 |  | n
Spaccarotella et al., (2022) | Assessment of Non-Invasive Measurements of Oxygen Saturation and Heart Rate with an Apple Smartwatch: Comparison with a Standard Pulse Oximeter | 2022 | 2022-03-08 | Nellcor Portable SpO2 Patient Monitoring System, PM10N (Medtronic, Minneapolis, MN, USA) | Series 6 | SpO2 | Of these 257 subjects, 56 were healthy controls, 60 were patients with lung disease, and 141 were patients with cardiovascular disease | y | y | 257 | 168 | 89 |  | n
Walzel et al., (2023) | Evaluation of Leading Smartwatches for the Detection of Hypoxemia: Comparison to Reference Oximeter. | 2023 | 2023-11-14 | Masimo Radical-7 (Masimo Corporation, Irvine, CA, USA) | Series 8 | SpO2 | Healthy young adults | y | n | 18 | 14 | 4 |  | n

## EE
First author (year) | title | publication_year | publication_date | criterion | aw_model | metric | Population cohort | healthy | clinical_pop | total_participants | Males | Females | Procedure | free_living? | at_rest?
Düking et al., (2020) | Wrist-Worn Wearables for Monitoring Heart Rate and Energy Expenditure While Sitting or Performing Light-to-Vigorous Physical Activity: Validation Study | 2020 | 2020-05-06 | Portable indirect calorimetry | Series 4 | energy expenditure | Healthy adults | y | n | 25 | 11 | 14 | sitting, walking and running at different speeds (1.1 m/s, 1.9 m/s, 2.7 m/s, 3.6 m/s, and 4.1 m/s) for 5 minutes, interspersed with 5 minutes of standing still. All participants also performed 6 ~30-m sprints involving multiple changes in direction /// two lab visits, 3 days between visits, 2 different wearables tested on each occasion, sitting, walking and running at different speeds (1.1 m/s, 1.9 m/s, 2.7 m/s, 3.6 m/s, and 4.1 m/s) for 5 minutes, interspersed with 5 minutes of standing still. All participants also performed 6 ~30-m sprints involving multiple changes in direction (ranging from 10° to 180°) | n | y
Falter et al., (2019) | Accuracy of apple watch measurements for heart rate and energy expenditure in patients with cardiovascular disease: Cross-sectional study | 2019 | 2019-03-19 | Indirect calorimetry | Sport (first-generation) 42 mm | energy expenditure | Patients with cardiovascular disease // All patients scheduled for a cardiopulmonary exercise test (CPET) as part of their cardiovascular rehabilitation program were consecutively included | n | y | 40 | 32 | 8 | Graded CPET on a cycle ergometer // The CPET started with 1 minute of seated rest. The exercise then started at 20 watts and load was increased with 20 W/min | n | n
Mihyun Lee et al., (2018) | Accuracy of swimming wearable watches for estimating energy expenditure. | 2018 | 2018-06-11 | Portable indirect calorimetry | Series 2 | energy expenditure | Healthy adults | y | n | 78 | 40 | 38 | participants swam for 4 minutes each at 0.4m/s , 0,6, 0.8, 1.0, 1.2m/s | n | n
Nuss et al., (2019) | Assessment of Accuracy of Overall Energy Expenditure Measurements for the Fitbit Charge HR 2 and Apple Watch | 2019 | 2019-05-01 | Indirect calorimetry | not reported | energy expenditure | Young adults | y | n | 30 | 15 | 15 | The protocol began with 3 minutes of upright sitting in a chair, followed by 2 minutes of stand- ing on the treadmill. After this 5-minute period, the Bruce protocol was performed. 3 mins walking post test, then 2 min sitting. | n | y
Nuss et al., (2019) | Accuracy of Heart Rate and Energy Expenditure Estimations of Wrist-Worn and Arm-Worn Apple Watches | 2019 | 2019 | Indirect calorimetry | Series 1 | energy expenditure | Young adults | y | n | 30 | 15 | 15 | The exercise protocol started with the participant sitting in a chair for three minutes, followed by standing on the treadmill for two minutes as HR readings were taken from each device every minute. The lowest ECG reported HR from the rest period was recorded as the participant’s seated HR. After this ve-minute rest period, the participant started the Bruce protocol. // Research staff recorded HR estimations from each device and from the ECG at the end of each minute of the Bruce protocol. | n | y
Sun et al., (2020) | Validity of Apple Watch 6 and Polar A370 for monitoring energy expenditure while resting or performing light to vigorous physical activity | 2020 | 2020-07-20 | Portable indirect calorimetry | Series 6 | energy expenditure | Young physically active males | y | n | 11 | 11 | 0 | This study was a randomized crossover trial, in which participants were randomly allocated into Group A or Group B. In Group A, ground running was performed first and then treadmill running, whereas Group B started with treadmill running and then ground running. A 72-h interval between the treadmill and ground experiments was applied to eliminate the effects of fatigue. //// each participant sat in a chair for at least 5 min, remaining awake and quiet, until his heart rate stabilized, to establish a baseline resting state. Resting EE was measured for 6 min. The exercise protocol involved a sequence of different running levels: 6.4 km/h, 8.0 km/h, 9.6 km/h, and 11.2 km/h (Supplementary Fig. 1). Each speed was kept constant for 6 min running on the treadmill. After completing each level of running, the participants were asked to sit in a chair again until their heart rates and oxygen consumption returned to the resting state. The procedure of the ground test was identical to that of the treadmill test. Conical barrel and sound cues were used to maintain a constant running speed. Conical barrels were placed every 10 m on the indoor round runway and the intervals of prompt tone were 6.625, 4.500, 3.750, and 3.210 s for 6.4, 8.0, 9.6, and 11.2 km/h, respectively. | n | y
Uphill et al., (2024) | Validity of Apple Watch, Garmin Forerunner® 935 and GENEActiv for estimating energy expenditure during close quarter battle training in Special Forces soldiers | 2024 | 2024-03-18 | Portable indirect calorimetry | Series 5 | energy expenditure | Male soldiers | y | n | 34 | 34 | 0 | Close quarters combat training activity | n | n
Wallen et al., (2016) | Accuracy of heart rate watches: Implications for weight management | 2016 | 2016-05-27 | Portable indirect calorimetry | likely Series 0 although not reported | energy expenditure | Healthy adults | y | n | 22 | 11 | 11 | 1-hr proto- cols involving supine and seated rest, walking and running on a treadmill and cycling on an ergometer. /// Participants initially performed five-minute periods of supine, sit- ting and standing, respectively. Three stages of a Bruce graded treadmill exercise protocol were then undertaken followed by five minutes of seated rest. Participants then completed six, three- minute stages of a 25-watt step test (commencing at 25 W) on a cycle ergometer followed by a final five minutes of seated rest. | n | y

## ECG waveform
First author (year) | title | publication_year | publication_date | criterion | aw_model | metric | Population cohort | healthy | clinical_pop | total_participants | Males | Females | Procedure | free_living?
Behzadi et al., (2020) | Feasibility and Reliability of SmartWatch to Obtain 3-Lead Electrocardiogram Recordings | 2020 | 2020-09-07 | 12-lead ECG | Series 4 | ecg waveform | Patients in sinus rhythm | n | y | 100 | 62 | 38 |  | n
Buelga Suárez et al., (2022) | Smartwatch ECG Tracing and Ischemic Heart Disease: ACS Watch Study | 2022 | 2022-11-21 | 12-lead ECG | Series 7 | ecg waveform | Patients with acute IHD | n | y | 25 | 18 | 7 |  | n
Harmon et al., (2022) | PERFORMANCE AND ACCURACY OF A SMART WATCH SINGLE-LEAD ECG: A PILOT STUDY | 2022 | May 2022 | 12-lead ECG | Series 5 | ecg waveform | Mayo Clinic ECG lab patients >18yrs, 6 patients were in AFib, 6 had pacemakers | n | y | 74 | 41 | 33 |  | n
Khushhal et al., (2025) | Accuracy of Apple Watch in Measuring 30-Second Resting Electrocardiography in Patients with Cardiac Diseases and Comorbidity: An Observational Cross-Sectional Study. | 2025 | 2025-01-30 | 12-lead ECG | Series 8 | ecg waveform | 34 cardiac patients with sinus rhythm, and 14 cardiac patients with irregular HR rhythm. Patients with chronic diseases, including 14 diabetics, 34 hypertension patients, and 16 dyslipidemia patients | n | y | 112 | 112 | 0 | simultaneous ECG and Apple Watch measurement | n
Klier et al., (2023) | Diagnostic Accuracy of Single-Lead Electrocardiograms Using the Kardia Mobile App and the Apple Watch 4: Validation Study. | 2023 | 2023-11-23 | 12-lead ECG | Series 4 | ecg waveform | Healthy adults | y | n | 81 | 58 | 23 |  | n
Saghir et al., (2020) | A comparison of manual electrocardiographic interval and waveform analysis in lead 1 of 12-lead ECG and Apple Watch ECG: A validation study | 2020 | 2020-08-26 | 12-lead ECG | Series 4 | ecg waveform | Healthy adults | y | n | 43 | 15 | 28 |  | n
Strik et al., (2020) | Validating QT-Interval measurement using the apple watch ecg to enable remote monitoring during the COVID-19 pandemic | 2020 | 2020-07-28 | 12-lead ECG | Series 4 | ecg waveform | 35% had diagnosed cardiac disease. | n | y | 100 | 59 | 41 |  | n

## Wheelchair push
First author (year) | title | publication_year | publication_date | criterion | aw_model | metric | Population cohort | healthy | clinical_pop | total_participants | Males | Females | free_living?
Benning et al., (2021) | Measurement Performance of Activity Measurements with Newer Generation of Apple Watch in Wheelchair Users with Spinal Cord Injury | 2021 | 2021-12-02 | Manual counting | Series 4 | wheelchair push count | People with subacute spinal cord injuries. Thirteen participants were diagnosed with paraplegia and two participants were diagnosed with tetraplegia. | n | y | 15 | 12 | 3 | n
Glasheen et al., (2021) | Accuracy of Apple Watch fitness tracker for wheelchair use varies according to movement frequency and task | 2020 | 23/03/2020 | Manual counting | Series 1 | wheelchair push count | 15 Wheelchair users with disabilities, and 15 able-bodied individuals | y | y | 30 | 18 | 12 | n
Karinharju et al., (2019) | Validity of the Apple Watch® for monitoring push counts in people using manual wheelchairs | 2019 | 2019-02-27 | Manual counting | Series 1 | wheelchair push count | Wheelchair users who are moderately physically active | n | y | 26 | 20 | 6 | n

## Step count
First author (year) | title | publication_year | publication_date | criterion | aw_model | metric | Population cohort | healthy | clinical_pop | total_participants | Males | Females | Procedure | free_living?
Bunn et al., (2018) | Assessment of step accuracy using the Consumer Technology Association standard | 2018 | 2018-06-29 | Manual counting | Series 1 | steps | Adults at low-mod risk of cardiovascular disease | y | n | 20 | 10 | 10 | Self-paced walking and running on treadmill for 5 mins each | n
Veerabhadrappa et al., (2018) | Tracking Steps on Apple Watch at Different Walking Speeds | 2018 | 2018-04-09 | Manual counting | not reported | steps | Adults | y | n | 71 | 47 | 24 | Each participant completed a treadmill protocol with four different speeds [slow (1.9mph), moderate (3.0mph), brisk (4.0mph), and jogging (5.2mph)] with 6 min duration/ speed. | n
Wallen et al., (2016) | Accuracy of heart rate watches: Implications for weight management | 2016 | 2016-05-27 | Manual counting | not reported | steps | Healthy adults | y | n | 21 | 11 | 11 | participants were videoed while doing a Bruce treadmill test and the steps count was obtained via video | n

## Sleep
First author (year) | title | publication_year | publication_date | criterion | aw_model | metric | Population cohort | healthy | clinical_pop | total_participants | Males | Females | Procedure | free_living?
Apple (2023) | Estimating sleep stages from Apple Watch | 2023 | september_2023 | PSG, home EEG | not reported | sleep | healthy, and those with sleep conditions | y | y | 166 | 88 | 78 | in-lab PSG and at-home EEG-based recording | y
Lee et al., (2023) | Accuracy of 11 Wearable, Nearable, and Airable Consumer Sleep Trackers: Prospective Multicenter Validation Study | 2023 | 2023-11-02 | PSG | Series 8 | sleep | Individuals recruited from hospital and clinic settings, all had presence of subjective sleep discomfort | n | y | 26 | not reported | not reported | Each group included a combination of noninterfering CSTs. Specifically, the nearables Google Nest Hub 2 and Amazon Halo Rise, which use similar radar sensors to detect motion, were allocated to different groups. In the case of wearables, participants were allowed to simultaneously wear a maximum of 2 watch devices, which are a type of wearable, with 1 on each wrist. One night of sleep recorded at a lab | n
Robbins et al., (2024) | Accuracy of Three Commercial Wearable Devices for Sleep Tracking in Healthy Adults. | 2024 | 2024-10-10 | PSG | Series 8 | sleep | Healthy adults | y | n | 35 | 15 | 20 |  | n

## VO2 max
First author (year) | title | publication_year | publication_date | criterion | aw_model | metric | Population cohort | healthy | clinical_pop | total_participants | Males | Females | Procedure | free_living?
Lambe et al., (2025) | Investigating the accuracy of VO2 max measurements from Apple Watch: A validation study | 2025 |  | indirect calorimetry | Series 5 to Series 9 | vo2max | healthy adults | y | n | 30 | 15 | 15 | lab-based indirect calorimetry on a treadmill, compared with Apple Watch estimates obtained in free-living setting | y

## Sleep apnea
First author (year) | title | publication_year | publication_date | criterion | aw_model | metric | Population cohort | healthy | clinical_pop | total_participants | Males | Females | Procedure | free_living?
Apple (2024) | Estimating Breathing Disturbances and Sleep Apnea Risk from Apple Watch | 2024 | september_2024 | home sleep apnea test | not reported | sleep apnea detection | healthy and those with sleep apnea | y | y | 1499 | 652 | 847 | This was a large clinical study, conducted to submit the feature for FDA approval | y

## Sound exposure
First author (year) | title | publication_year | publication_date | criterion | aw_model | metric | Population cohort | healthy | clinical_pop | total_participants | Males | Females | Procedure | free_living?
Fischer et al., (2022) | Are Smartwatches a Suitable Tool to Monitor Noise Exposure for Public Health Awareness and Otoprotection? | 2022 | 23/03/2022 | Class 1 sound level meter | Series 6 | sound exposure | na | na | na | na | na | na | Noise measured in various environments with sound level meter and with AW through the Noise app | n

## HTN notification
First author (year) | title | publication_year | publication_date | criterion | aw_model | metric | Population cohort | healthy | clinical_pop | total_participants | Males | Females | Procedure | free_living?
Apple (2025) | Hypertension Notification Feature on Apple Watch | 2025 | 2024-09-01 | OMRON Evolv® Wireless Upper Arm Blood Pressure Monitor (Model BP7000, K162092). | not reported | HTN notification | healthy and those with HTN, and those just with HTN risk factors | y | y | 2229 | 1268 | 961 | Performance was determined based on average blood pressure readings taken twice daily at home over 30 days using the reference device, the OMRON Evolv® Wireless Upper Arm Blood Pressure Monitor (Model BP7000, K162092). // Of the 2229 participants, 1863 met usable data requirements and contributed to the primary endpoint analyses. Those not included in the performance evaluation lacked sufficient Apple Watch or reference data or had major protocol violations. In the Notification Analysis set, 585 participants exhibited hypertension, while 1278 did not. | y

## HRV
First author (year) | title | publication_year | publication_date | criterion | aw_model | metric | Population cohort | healthy | clinical_pop | total_participants | Males | Females | Procedure | free_living?
O'Grady et al., (2024) | The Validity of Apple Watch Series 9 and Ultra 2 for Serial Measurements of Heart Rate Variability and Resting Heart Rate. | 2024 | 2024-09-26 | Polar H10 + Kubios | Series 9 and Ultra 2 | hrv | healthy adults | y | n | 39 | 17 | 22 |  | y

## 6MWD
First author (year) | title | publication_year | publication_date | criterion | aw_model | metric | Population cohort | healthy | clinical_pop | total_participants | Males | Females | Procedure | free_living?
Apple (2021) | Using Apple Watch to Estimate Six Minute Walk Distance | 2021 | may_2021 | supervised 6MWT | not reported | 6MWD | Inidividuals aged 65 or older across a broad spectrum, including those with arthritis, copd, diabetes, coronary artery disease, and using canes/walkers or no assistive device | y | y | 449 | 184 | 255 | adults ages 65 and older who live either in the community or in independent living housing // supervised 6MWTs consistent with published guidelines on a linear and flat “out and back” course ranging from 15 to 30 meters in length // Participants completed up to 5 reference 6MWT tests | y